Troubleshooting the Freedom Processor

# Text boxes

Maintenance Check

Confirm:

- Transmitting cable is properly connected

- Cable is not broken or cracked

- Transmitting coil is on the child’s head

Also do the following:

- Confirm that the Freedom BTE is ‘ON’

- If no sound, turn the Freedom BTE ‘OFF’ for about 3 seconds and then ‘ON’

- Confirm the child is using the recommended programme setting

- Identify whether HELP messages are displayed on the LCD

- H3 = Coil Error
Check the coil is fully inserted into the socket
Check the coil is correctly located over the implant
Check for cable damage
Change cable coil

H3

- H2 = Battery Low
Change all the batteries

H2

- H1 = Battery Flat
Change all the batteries
Ensure all batteries are correctly inserted (+ve down)
Change the battery holder (with new batteries)

H1

Helper Messages

H5

- H5 = MAP corrupt
Try another program location
Contact clinic

H4

- H4 = sound/stimulation error
Turn processor “Off” then “On”
Check/Change microphone protector cover
Check/Change coil
Check sensitivity is above “0”
Check with lapel microphone
Contact clinic to change the processor